## Slide 1
2021
Axin

## Slide 2
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Take home message

## Slide 3
AxIN®
What is it ?
products
A customized SR tooth in zirconia with an ingenious Angulated Access performed in Simeda® CMC
Notes: 1 minute 50

## Slide 4
specific CAD CAM machining compartment
Crown in Zirconium milled by Simeda®
Anthogyr spherical key
AxIN®
What is it ?
AxIN base + screw
Notes: The AxIN solution consists of a titanium base and a screw, where when the screw is tightened some wings open and block the crown made by CADCAM Simeda, which has a cavity with a specific machining compartment.

## Slide 5
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Take home message

## Slide 6
AxIN® Indications
For BL and TL, ALL sectors:
  Anterior area : Aesthetic +++ and reduced maintenance
    Pre-molar : Aesthetic + and reduced maintenance
  Molar: Aesthetic and reduced maintenance Accessibility to screwing Available for Axiom BL Ø 4.0 / 5.0 Ht 2.5 and for Axiom TL ht neck 2.5 and 3.5 CONTRA-INDICATED in Ht 1.5 mm on Axiom BL and Axiom TL  Anticipate the choice of implant and restoration before surgery
biologically safe
Single unit restorations

## Slide 7
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Take home message

## Slide 8
products
AxIN® Benefits
Biological safety
Promote screw-retained tooth
  No sealing cement
  Less risk of peri-implantitis
Mechanical assembly of the titanium base
  No glue
  Preservation of soft tissues
Titanium and zirconia material
  Excellent biocompatibility

## Slide 9
products
AxIN® Benefits
Ergonomic solution
Single-piece tooth
Built in screw
  No handling of small screw in mouth
Angulated Access, free orientation between 0 and 25°
  Oriented channels facilitate tightening the screws in the mouth
  No implant indexation needed
  Reduced and easy maintenance
  Removable titanium base
  Angulated Access, free orientation between 0 and 25°
  Oriented channels to limit the risk of chipping
Easy handling for Dentist and Lab
Notes: Easy handling : remodelling without oven?
Built in screw (visual) no loss or risk of error
In lab, handling is easier than a titanium base thanks to the lab base (regulated friction) false supple gingiva, buried prosthesis => « spring » effect => rises => difficult
Very easy : impression directly on implant, same material, spherical screwdriver, extractor (less equipment/material?)
Receive from prosthetist the finished tooth
Single position : (customized) indexation => whatever the trilobe’s position in the mouth
Peculiarity : the screw is already part of the base => no transport, simplicity, handling comfort, implemented/established?

## Slide 10
products
AxIN® Benefits
Closer to the natural tooth
Screw inserted through the bottom part
  Narrow channel diameter
  Preserved mechanical integrity
  Ceramic support
AxIN®
Market standards
Minimum height: 4.9mm
Minimum diameter: 4.5mm
Notes: Positioning of the implant in the root orientation.
No need of placing the apex on palatin side.

## Slide 11
Aesthetic excellence
Angulated Access, free between 0 to 25°
Narrow screw channel (2.0 mm)
From provisional to definitive
Temporary angulated solution to pre-shape soft tissues
AxIN® Benefits
products
Notes: ML, Z, T
Customized emergence profile
Haut de gamme, sur mesure Simeda, anatomique, précisions, logement spécifique
Zirconia tooth
(CAD CAM tooth)

## Slide 12
products
AxIN® Benefits
Aesthetic & efficient prosthetic innovation
Screw-retained anatomical tooth
  Biological safety : no cement, no glue
  Aesthetics : narrow channel Angulated Access
  User friendly : integrated screw and Angulated Access
  Choice of zirconia
  For BL and TL, all sectors
  Serenity® guarantee
New vision
Closer to the natural tooth
Simplified treatment
Good to know* !
2/3 of cases are placed in posterior areas.
Notes: Accessibility to innovation
• Immediate loading solutions
with angled provisional components
INNOVATION january 2019
Zirconia Sina Z, T, ML

## Slide 13
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Take home message

## Slide 14
BL DEFINITIVE COMPONENTS
Hg 1.5
Hg 2.5
Hg 1.5
Hg 2.5
Abutment Extractor
Height identification on base and screw
Monogram on three lobed
BASE Ø 4.0
BASE Ø 5.0
Definitive M1.6 screw
Identification with lines for each height
AxIN® Components
Ht 1.5 : contra-indicated in molar sector
Ball wrench key

## Slide 15
BL LAB COMPONENTS
Lab base
Axiom BL Ø 4.0
Delivered with engaged screw
M1.6 BL AA lab screw
Anodized blue
1 per gingival Ht
Hg Identification
Ball wrench key
AxIN® manipulation key For axiom® Multi Level®
Hg 1.5
Hg 2.5
Lab base
Axiom BL Ø 5.0
Delivered with engaged screw
Height identification on base and screw
Very light screwing
< 5 N.cm
AxIN® Components

## Slide 16
TL DEFINITIVE COMPONENTS
TL N
TL R
Common R/N
Ball wrench key
Monogram on three lobed
TL BASE
Definitive M1.6 screw
Bicolor black/titanium
AxIN® Components
Contra-indicated in molar sector on Axiom TL implant N and R 1.5 ht neck

## Slide 17
TL LAB COMPONENTS
TL lab base
Delivered with engaged screw
M1.6 BL AA lab screw
Anodized blue
Ball wrench key
AxIN® manipulation key For axiom® Multi Level®
TL N
TL R
Very light screwing
< 5 N.cm
AxIN® Components
Common R/N

## Slide 18
AxIN® Components
AxIN® manipulation key for axiom® Multi Level®
Out of mouth use : for lab and practitioner
  Disassemble the crown from the final base.
  Optional: easily assemble the crown on the definitive base..
AxIN BL/TL base is held at the level of the BL/TL tri-lobe.

## Slide 19
PROVISIONNAL COMPONENTS
Temporary coping 10, 15 or 25°
AxIN® Components
Optional : protective cap Ø 4.0, 4.8, 5.0
Used at the dental office, to cover the base in mouth

## Slide 20
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Competitors
  Take home message

## Slide 21
Workflow temporary tooth
At the lab
At the practitioner’s office (2 possibilities)
AxIN® Provisional workflow
Notes: Ht mini = 4.4 con deflexión de tornillos + 0.5 espesor de prótesis arriba
Dia en la plataforma N o 4.0 = 4.3 a 0.5 mm para respetar el grosor mini - bajo exocad / DW / 3shape

## Slide 22
AT THE LAB
AxIN® Provisional workflow
Choice of the provisory abtmt
Ht, Ø, angulation
Assembly of the provisory tooth
Provisory coping +
Lab base
Ti-base and its screw – 1 bag
Coping – 1 bag
Ti-base with a three lobed indexation
Making of the provisory on model
With a market tooth
Send to the practitioner

## Slide 23
AT THE PRACTITIONER’S OFFICE – 1st possibility - Made on definitive Ti-base
AxIN® Provisional workflow
Choice of the provisory abtmt
Ht, Ø, angulation
Assembly of the provisory tooth
Provisory coping + Ti-base
Ti-base and its screw – 1 bag
Coping – 1 bag
Ti-base with a three lobed indexation
Making of the provisory in and outside the mouth
With a market tooth

## Slide 24
AT THE PRACTITIONER’S OFFICE – 2nd possibility – Made on lab base
AxIN® Provisional workflow
Choice of the provisory abtmt
Ht, Ø, angulation
Assembly of the provisory tooth
Provisory coping + lab base
Ti-base and its screw – 1 bag
Coping – 1 bag
Ti-base with a three lobed indexation
Making of the provisory outside the mouth
With a market tooth
+ Protected gingiva
Optional :
Protective cap on base in mouth

## Slide 25
AxIN – PROVISIONAL – Recommendations:
It is advised to disinfect the provisional before placing it in mouth.
index the implant correctly during the surgery
Do an x-ray to control
Close the screw shaft
AxIN® Provisional workflow

## Slide 26
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  Manipulation
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Competitors
  Take home message

## Slide 27
LAB WORKFLOW
CAD design (exocat, DW, 3shape)
or sent model / wax-up

web order*
Reception of crown and final base, unassembled
Crown finishing on laboratory basis
Assembling the crown in its correct position on the final base with its screw before sending it to the clinic
AxIN® Definitive workflow
Reminder: Before the 1st case:
download the new implant library with AxIN
Done by Mersch
Reco : delivered not mounted on the model
* AxIN contraindicated in Ht 1.5 in molar sector => choose base plate Ht 2.5, if case BL
Notes: Deliver the finished AxIN tooth (e.g. in its Simeda box) not mounted on the model (otherwise it will be disassembled in 3 steps and reassembled by the dentist).

## Slide 28
AxIN® Definitive workflow
Cut back
Homothetical
0°
15°
Design options
Homothetical / Reduction and ordering of sina Z/T/ML/XT-T
Cut back » and ordering of sina ML

## Slide 29
AxIN® Definitive workflow
Ceramisation of the crown
Lab Ti-base are mandatory - screwed at 5 N.cm
The crown platform must be clean of ceramic (glazure) to have a perfect fit on Ti-base, a correct deployment of the wings

## Slide 30
PRACTITIONER WORKFLOW
Reception of the tooth
Placement on the implant
Screwing with ball wrench key
Check with an X ray
AxIN® Definitive workflow
OK
NOT
OK

## Slide 31
AxIN – DEFINITIVE CROWN
Can be performed by laboratories with and without scanner
Angulated access up to 25º
Minimum height of the prosthesis above the base: 4.9 mm
Minimum diameter of the prosthesis 4.5 mm
1.5 height contra-indicated in the molar sector
AxIN® Definitive workflow

## Slide 32
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  Manipulation
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Take home message

## Slide 33
Case: Illustration of the gingival development and the preparation of a temporary screw retained prosthesis
Dr Margossian - TITANE – Revue Française Vol 16 / N° 3 / Sep 2019
AxIN®clinical case
Notes: a 32-year-old patient, lost her central incisor in an accident. A bone reconstruction has been made with a GBR at the same time as the placement of the implant.

## Slide 34
AxIN® clinical case

## Slide 35
AxIN® clinical case

## Slide 36
AxIN® Clinical Experience – EU Review
Notes: Angulated Access, free between 0 to 25°
Narrow screw channel (2.0 mm)
Temporary angulated solution to pre-shape soft tissues : 10°, 15°, 25°
Conclusion : Continuous innovation in the field of implantology allowed the optimization of the results biological, aesthetic and functional. These latest developments
caused the obsolescence of cemented solutions.
Beyond a simple prosthetic evolution,it's a paradigm shift in positioning
anterior implant that is set up in the service of anatomy and biological

## Slide 37
AxIN® Clinical Experience
Dr Margossian (Marseille – FR) : “Change of paradigm in anterior implant positioning” TITANE – French review Vol 16 / N° 3 / Sep 2019
With straight channel
A : respects the anatomy, but channel in the free edge.
B : an ideal cingulate emergence, but apex of the implant close to the vestibular wall
C : apex of the implant close to the palatal wall.
With angulated channel
Angulated Access allows an ideal positioning of the implant in the bone anatomy, while allowing an exit into the cingulum in order to preserve the vestibular aesthetics.
Illustration of the key differences in the design of two options of axis adjustment. The AxIN system on the left allows, among others, a preservation of material that ensures more room for layering in the vestibular area, and. better esthetic results

## Slide 38
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Take home message

## Slide 39
LAB
PRACTITIONER’S OFFICE
AxIN lab kit or the AxIN key + lab base + ball wrench
« order on e.shop or on Anthogyr Simeda weborder “
equipped with the ball wrench ? + order the BL slimmed down extractor abutment (highly recommended) + AxIN key (advisable +++)
=> « order on e.shop ”
AxIN® ancillaries
Instruments

## Slide 40
Summary
AxIN
  What is AxIN?
  Indications
  Benefits
  Components
  AxIN provisional workflow
  AxIN definitive workflow
  Clinical case
  Ancillaries for lab and practitioner
  Take home message

## Slide 41
AxIN®
TAKE HOME MESSAGE
A customized SR tooth in zirconia with an ingenious Angulated Access performed in Simeda® CMC
Closer to natural tooth
Biological safety : no cement, no glue
User friendly for the practitioner and the dental technician
BL & TL consistency, all sector (except Ht 1.5 in molar area)

## Slide 42
Choose your new restoration